# SaxStat Prototypes v02 and v03

2025.7.30

## Overview

This document provides a detailed technical breakdown of the SaxStat prototypes v02 and v03, focusing on hardware components, firmware structure, and GUI implementation. It aims to document the system for future development, including macroelectrode applications and potential UME enhancements, with PCB design considerations.

## Hardware Details

### Prototype v02 Hardware

- Microcontroller: ESP32-DevKitC (dual-core Tensilica LX6, 520 KB SRAM, 4 MB flash), with 2x19 male headers.

- Digital-to-Analog Converter (DAC): AD5761, 16-bit bipolar DAC, TSSOP-16, providing voltage output from -10 V to +10 V (with ±5 V supply).

- Analog-to-Digital Conversion: ESP32 internal ADC via GPIO34, 12-bit resolution, 0-3.3 V range, with level-shifting to prevent negative signals.

- Transimpedance Amplifier (TIA): LMC6484 op-amp configured with jumper-selectable resistors: 40 kΩ (R4) or 1 MΩ (R8), targeting 25 nA/V gain for macroelectrodes (e.g., 12.5 µA peak with 2 mM ferrocene-dimethanol).

- Level-Shifting: Vref=1.0 V via a resistor voltage divider at op-amp D of LMC6484, shifting TIA output to ensure non-negative input to GPIO34.

- Power Supply: Single 5 V via USB, with ICL7660 charge pump (U7) generating ±5 V. ADR4525 (U8) provides a 2.5 V reference.

- Passives: R9/R10 (4.7 kΩ I2C pull-ups), C1/C2 (10 µF for ICL7660), C3-C6 (0.1 µF decoupling), jumper headers for R4/R8 selection.

- Connections: SPI for AD5761, 3-pin terminal block (J2) for WE, CE, RE electrodes.

### Prototype v03 Hardware

- Microcontroller: ESP32-DevKitC, unchanged.

- DAC: AD5761, with a low-pass filter (e.g., 10 kΩ resistor and 0.1 µF capacitor) added after the output (U3 Pin 1, VOUT) to reduce noise.

- ADC: ADS1115, 16-bit resolution, 4-channel, 4.096 V reference, with a low-pass filter (e.g., 10 kΩ resistor and 0.1 µF capacitor) added before the input (U4 Pin 7, AIN0) to smooth TIA output.

- TIA: LMC6484 with tested resistors: 10 kΩ and 1 MΩ, with experimental use of 10 MΩ for UME testing.

- Power Supply: Same ±5 V setup with added 0.1 µF bypass capacitors near the TIA and filters.

- Connections: SPI for AD5761, I2C for ADS1115 (address 0x48 with ADDR to GND), with filtered analog paths.

### Hardware Challenges

- v02: GPIO34’s 12-bit resolution and noise limited accuracy.

- v03: Noise with 1 MΩ/10 MΩ TIAs persisted despite filters, possibly due to LMC6484 noise floor (~1 nA) or shielding. Filters improved stability but didn’t resolve UME distortions at ~1 nA.

## Firmware Structure

### SaxStat_V02_GUI_Test.ino

- Purpose: Drives the AD5761 DAC and reads ESP32 GPIO34 ADC, interacting with a Python GUI for CV.

- Libraries: Uses SPI.h for AD5761 communication.

- Key Definitions:

  - SCK_PIN (18), MOSI_PIN (23), CS_PIN (5) for SPI.

  - ADC_PIN (34) for ESP32 ADC input.

  - AD5761_CLOCK_RATE (10 MHz), control register for ±3 V range.

  - DAC codes: MIN_DAC_CODE (15934) for -1.5 V, MAX_DAC_CODE (49602) for +1.5 V.

- Key Functions:

  - ad5761_write(uint8_t reg_addr_cmd, uint16_t reg_data): Writes to AD5761 via SPI with 3-byte protocol.

  - mapVoltageToDAC(float voltage): Maps -1.5 to 1.5 V to DAC codes.

  - setup(): Initializes serial (115200 baud), SPI, sets AD5761 to 0 V, 5-second delay for stability.

  - loop(): Parses serial commands, manages CV ramp with running state.

- Command Handling:

  - START:<start_v>:<end_v>:<scan_rate>:<cycles>: Sets parameters, calculates halfCycleTimeMs, stepSize, delayMs, initiates ramp with 50-sample initial skip.

  - STOP: Halts ramp, resets to 0 V.

  - CALIBRATE: Sends 100 ADC samples.

  - RESET_GPIO34:0: Resets DAC to 0 V.

  - READ_INITIAL: Reports initial ADC voltage.

- Output: Raw ADC values (0-4095), setup messages (e.g., "minValue (DAC): ..."), "CV complete."

- Challenges: Limited by GPIO34 resolution, transient data handling.

### SaxStat_V03_GUI_Test.ino

- Purpose: Extends v02 with ADS1115 driver for improved resolution.

- Libraries: Adds Wire.h for I2C.

- Key Definitions: Retains AD5761 pins, adds ADS1115_ADDRESS (0x48).

- Key Functions:

  - setup(): Initializes I2C, configures ADS1115 (gain=1, 128 SPS).

  - read_ads1115(): Reads ADS1115 via I2C, accounts for filter delay.

  - loop(): Adapts v02 logic for ADS1115, retains command structure.

- Command Handling: Same as v02, with ADS1115 data.

- Output: Raw ADS1115 values (0-32767), inconsistent "START_CONFIRMED" in UME tests.

- Challenges: Filter delay affects timing, assumes 10 kΩ TIA.

## GUI Structure

### SaxStat_GUI_V02.py

- Framework: PyQt5 for UI, pyqtgraph for plotting.

- Key Components:

  - CyclicVoltammetryGUI: Left panel (inputs: start/end voltage, scan rate, cycles, offset), right panel (voltage-time, CV graphs).

  - start_acquisition(): Sends START, starts read thread with 2-second timeout.

  - read_serial_data(): Reads GPIO34 values, converts to current with 30 kΩ TIA, applies 10-point smoothing.

  - update_plots(): Updates graphs with ±10 µA Y-range.

- Data Processing:

  - Current = (1e6 * (2 * V_ref - V_out - V_applied) / 30,000) µA.

  - Offset calibration averages 100 samples over 1 second.

- Features: COM port refresh, connect/disconnect, save to CSV.

- Variables: resistance = 30000 (Ω), v_ref = 1.0 (V).

### SaxStat_GUI_V03.py

- Framework: Same as v02, with minor tweaks.

- Key Components:

  - CyclicVoltammetryGUI: Updated offset label to "Offset Current (nA)", adjusted read_serial_data() for ADS1115 (0-32767), 10 kΩ/1 MΩ TIA support.

  - start_acquisition(): Unchanged.

  - read_serial_data(): Scales current to nA, accounts for filter delay.

  - update_plots(): Adjusted for nA range (±100 nA).

- Data Processing:

  - Current = 1e6 * (2 * V_ref - V_out - V_applied) / resistance uA, where resistance is 10,000 or 1,000,000.

- Changes: Minimal, focused on TIA flexibility and ADS1115 integration.

## PCB Design Considerations

- Component Layout:

  - Place ESP32-DevKitC centrally, LMC6484/AD5761/ADS1115 near J2 (electrode terminal), ICL7660 near ESP32, ADR4525 near AD5761.

- Power Supply:

  - Route ±5 V from ICL7660 with 0.1 µF bypass capacitors. Add 10 µF capacitors (C1, C2) across charge pump outputs.

  - Include a 2.5 V reference line from ADR4525 with 0.1 µF decoupling (C6).

- Analog Paths:

  - Use 10 mil traces for V_ramp (AD5761 VOUT) and OpC_OUT (LMC6484 OUT_C), with low-pass filters (10 kΩ, 0.1 µF) post-DAC and pre-ADC.

  - Shield analog traces with a ground plane.

- Digital Paths:

  - Route SPI (SCK, MOSI, CS) and I2C (SDA, SCL) with 4.7 kΩ pull-ups near ADS1115.

- Grounding:

  - Implement a star-ground point near ICL7660, connecting all GND pins to minimize loops.

- TIA Configuration:

  - Design jumper headers or DIP switches for 10 kΩ, 1 MΩ, 10 MΩ TIA resistors.

- Testing Points:

  - Add test pads for V_ramp, OpC_OUT, VREF, and ADC input.

- Fabrication:

  - Use EasyEDA, export Gerber files, order 5 PCBs from JLCPCB (2-layer, 1.6 mm, ~$5).

- Enhancements:

  - Consider TS5A3160 switches for automated TIA selection, ESP32-WROOM-32E-N8 for compactness.

## Conclusion

This documentation details v02’s ESP32/GPIO34 setup and v03’s ADS1115 with filters, capturing macroelectrode success and UME challenges. It provides a foundation for PCB design.