# SaxStat PCB v04 Specification

## Overview

The SaxStat PCB v04 is a low-cost, compact potentiostat designed for DIYers, electrochemical sensor developers, and educators. It is compatible with the ESP32-DevKitC-V4 (38-pin) development board, featuring female pin headers for easy integration. The PCB supports three-electrode electrochemical experiments (working electrode [WE], counter electrode [CE], reference electrode [RE]) for screen-printed electrodes, delivering reliable performance for cyclic voltammetry (CV) and other electrochemical techniques. The design prioritizes affordability, simplicity, and modularity, with a target sale price of ~$150.

## Target Audience

- DIYers: Enthusiasts building custom electrochemical sensors or IoT projects.

- Electrochemical Sensor Developers: Researchers prototyping sensors for environmental, medical, or industrial applications.

- Educators: Instructors teaching electrochemistry or embedded systems in academic settings.

## Key Features

- ESP32-DevKitC-V4 Compatibility: Female pin headers for direct plug-in of the 38-pin ESP32-DevKitC-V4 board, ensuring easy setup and prototyping.

- Three-Electrode System: Supports WE, CE, and RE for precise electrochemical measurements (macroelectrode focus, µA currents).

- Low-Cost Design: Total production cost ~$100 for 10 PCBs and components, leveraging JLCPCB manufacturing and assembly.

- Compact Form Factor: PCB dimensions optimized for breadboard compatibility and small enclosures (~50 mm x 50 mm, 2-layer FR4).

- User-Friendly Interface: PyQt5-based GUI for real-time data acquisition, plotting, and parameter control (e.g., scan rate, voltage range).

- Open-Source Firmware: Arduino-based firmware for ESP32, supporting CV and customizable electrochemical experiments.

## Hardware Specifications

### Microcontroller

- Module: ESP32-DevKitC-V4 with ESP32-WROOM-32E (4 MB flash, dual-core, 240 MHz, Wi-Fi, Bluetooth).

- Interface: Female pin headers (2x19, 38-pin layout, 2.54 mm pitch) for direct ESP32-DevKitC-V4 plug-in.

- GPIO Usage:

  - SPI: SCK (GPIO18), MOSI (GPIO23), CS (GPIO5) for DAC.

  - I2C: SDA (GPIO21), SCL (GPIO22) for ADC.

  - GPIO2, GPIO4 for analog switch control (TIA resistor selection).

  - Avoid D0, D1, D2, D3, CMD, CLK (SPI flash pins) to prevent conflicts, as per ESP32-DevKitC-V4 documentation.

- Power: USB-powered via ESP32-DevKitC-V4 (5 V, regulated to 3.3 V onboard).

### Digital-to-Analog Converter (DAC)

- Component: AD5761 (16-bit, ±3 V output, SPI interface).

- Voltage Range: -2.5 V to +2.5 V (configurable via firmware).

- Output: V_ramp to TIA input via low-pass filter (10 kΩ, 0.1 µF) for noise reduction.

- Reference: 2.5 V from ADR4525 for high precision.

### Analog-to-Digital Converter (ADC)

- Component: ADS1115 (16-bit, I2C interface, 4.096 V reference).

- Input: Measures TIA output (OpC_OUT) via low-pass filter (10 kΩ, 0.1 µF).

- Reference: 2.5 V from ADR4525, with 1.0 V offset (ADR510) to ensure non-negative ADC input.

### Transimpedance Amplifier (TIA)

- Component: LMC6484 (quad op-amp, rail-to-rail, 1.5 MHz bandwidth, ~1 nA noise floor, as used in v03 prototype).

- Configuration: Selectable feedback resistors (10 kΩ, 1 MΩ) via TS5A3160 analog switches for macroelectrode (µA) measurements.

- Feedback Capacitor: 10 pF to prevent oscillation, based on v03 testing.

- Note: LMP7721 (3 fA bias current) deferred to future versions for UME support (nA currents).

### Analog Switches

- Component: TS5A3160 (single-channel, 5 V tolerant, low on-resistance).

- Function: Automated selection of TIA feedback resistors (10 kΩ, 1 MΩ) via ESP32 GPIO (GPIO2, GPIO4).

- Control: Firmware-driven, with GUI option for manual override.

### Power Supply

- ICL7660: Generates ±3 V from 5 V USB input for op-amp and DAC/ADC references.

- Capacitors: 10 µF (C1, C2) for ICL7660, 0.1 µF for decoupling (C3–C6).

- Reference Voltage: ADR4525 (2.5 V, low-noise) for DAC and ADC.

- Grounding: Star-ground point near ICL7660 to minimize noise and ground loops.

### Electrochemical Interface

- Connector: 3-pin terminal block (J2) for WE, CE, RE.

- Supported Experiments: Cyclic voltammetry (CV), linear sweep voltammetry (LSV), with macroelectrode focus (e.g., 75 µM ferrocenemethanol, as in SweepStat).

- Current Range: 100 nA to 100 µA (LMC6484-based TIA, 10 kΩ to 1 MΩ feedback).

### PCB Specifications

- Layers: 2-layer FR4, 1.6 mm thickness.

- Dimensions: ~50 mm x 50 mm (fits JLCPCB $2 pricing for 5 PCBs).

- Trace Widths:

  - Power: 15–20 mil for ±3 V, 5 V, GND.

  - Analog: 10 mil for V_ramp, OpC_OUT.

  - Digital: 8–10 mil for SPI, I2C.

- Grounding: Bottom-layer ground plane with star-ground point and vias for top-layer GND pads.

- Test Points: V_ramp, OpC_OUT, VREF, ADC input (AIN0) for debugging.

- Connectors: Female pin headers (2x19, 38-pin, 2.54 mm pitch) for ESP32-DevKitC-V4, 3-pin terminal block (J2), jumper headers for manual TIA resistor selection.

- Silkscreen: Clear labels for J2 (WE, CE, RE), test points, and key components.

### Firmware and Software

- Firmware: Arduino-based (SaxStat_V03_GUI_Test.ino as baseline), supporting:

  - SPI control for AD5761 (voltage ramp generation).

  - I2C control for ADS1115 (current measurement).

  - GPIO control for TS5A3160 (TIA resistor switching).

  - CV parameters: scan rate (10–1000 mV/s), voltage range (-2.5 V to +2.5 V).

- GUI: PyQt5-based (SaxStat_GUI_V03.py as baseline), featuring:

  - Real-time plotting of i-E curves.

  - User inputs for scan rate, voltage range, and TIA gain.

  - Data export to CSV for analysis.

- OTA Updates: Optional over-the-air firmware updates via ESP32 Wi-Fi, inspired by ESPHome recommendations.

### Performance Metrics

- Resolution: 16-bit (AD5761 DAC, ADS1115 ADC).

- Voltage Range: -2.5 V to +2.5 V (configurable).

- Current Range: 100 nA to 100 µA (macroelectrode focus).

- Noise Floor: ~1 nA (LMC6484 TIA, sufficient for macroelectrodes; future LMP7721 for <10 pA).

- Scan Rate: 10–1000 mV/s for CV experiments.

- Accuracy: ±0.1% for voltage, ±1% for current (based on v03 prototype testing).

### Cost Breakdown

- PCB Manufacturing: ~$2 for 5 PCBs (2-layer, 50 mm x 50 mm, JLCPCB).

- Components:

  - LMC6484: ~$2 each.

  - AD5761: ~$5 each.

  - ADS1115: ~$3 each.

  - TS5A3160: ~$0.50 each (x2).

  - ICL7660: ~$1 each.

  - ADR4525: ~$3 each.

  - Passives (resistors, capacitors): ~$2 total.

  - Connectors (J2, headers): ~$2 total.

- Total per PCB: ~$9–$10 (components) + ~$0.40 (PCB) = ~$10.

- Total for 5 PCBs: ~$50 (excluding shipping, ~$10–$20).

### Design Constraints

- ESP32-DevKitC-V4 (38-pin):

  - Dimensions: 25.5 mm width, 2.54 mm pin pitch, 38 pins (2x19).

  - Avoid using D0, D1, D2, D3, CMD, CLK (SPI flash pins).

  - GPIO16, GPIO17 available (ESP32-WROOM-32E module).

- Breadboard Compatibility: Female headers ensure plug-and-play with ESP32-DevKitC-V4, but PCB width may limit full breadboard insertion (use one side or offset mounting).

- Noise Management: LMC6484’s ~1 nA noise floor limits UME applications; ground plane and low-pass filters (10 kΩ, 0.1 µF) mitigate noise for macroelectrodes.

- Power: USB-powered (5 V), no external power option for simplicity.

### Testing and Validation

- Dummy Cell: Test with 1 MΩ resistor (as in SweepStat Figure S9) to verify linear i-E curves.

- Electrochemical Test: CV with 75 µM ferrocenemethanol (macroelectrode, ~µA currents) to validate performance.

- Debugging: Test points for V_ramp, OpC_OUT, VREF, and ADC input to diagnose issues.

- Shielding: Optional Faraday cage for low-current tests, as recommended by SweepStat (Figure S8).

### Future Enhancements

- LMP7721 Integration: Upgrade TIA to LMP7721 (3 fA bias current) for UME support (nA currents) in v05.

- FSCV Support: Add fast-scan cyclic voltammetry (FSCV) capabilities with higher bandwidth op-amps.

- Enclosure: Design 3D-printed enclosure (Fusion 360/FreeCAD) for portability and shielding (Task 3, workplan).

### Timeline

- Schematic Design: July 31–August 6, 2025.

- PCB Layout: August 7–13, 2025.

- Gerber Generation and Order: August 14–15, 2025 (JLCPCB, 5–7 day delivery).

- Component Sourcing: August 16–22, 2025 (DigiKey, Mouser, LCSC).

- Assembly and Testing: August 23–28, 2025.

- Pilot Launch: September 1, 2025 (per workplan).

### References

- SaxStat Prototypes v02 & v03.docx: Hardware baseline and testing results.

- SaxStat Pilot Workplan.markdown: Cost, timeline, and component guidance.

- SweepStat (ed9b00893_si_001.pdf): Benchmark for low-cost design and testing protocols.

- ESP32-DevKitC-V4 Documentation: Pinout and GPIO restrictions.